## V1_V2_Comments Combined
VOTING DRAFT STANDARD. EL-V1M1 AND V2M2. JUNE 2015. RESPONSE TO COMMENTS
Vote tally V1M1: Affirmative 81; Affirmative with Comment 13; Negative with Comment 12, Abstain 2
Vote tally V2M2: Affirmative 87; Affirmative with Comment 8; Negative with Comment 10; Abstain 3
Disclaimer: The NELAC Institute (TNI) accepts no liability for the content of any comment made on a standard. Any views or opinions on a standard are solely those of the commenter and do not necessarily reflect those of TNI.
| Vote | Section | Comment | Disposition | Change/Reasoning
| Affirmative with Comment | V1M1 3.1 | The definition may need to be revisited in the future if TNI proceeds with recognizing non-governmental accreditation bodies as okay for the NELAP program. Alternatively, are there plans for a different Volume 1, Module 1 for a non-state environmental laboratory TNI accreditation system? | Persuasive | Definition changed to be more aligned with ISO 17011.
| Negative with Comment | V1M1 3.1 | This definition potentially excludes Non-Governmental ABs. There is a major effort underway in TNI to recognize NGABs to accredit to the TNI Standard. 3.1 Accreditation Body: The territorial, state or federal agencyorganization having responsibility and accountability for environmental laboratory accreditation and which grants accreditation under this program. | Persuasive | Definition changed to be more aligned with ISO 17011.
| Affirmative with Comment | V1M1 3.1 | Does not match the definition in V2M2. These should be consistent. | Persuasive | Definition changed to be more aligned with ISO 17011.
| Negative with Comment | V1M1 3.12 | Editorial and to add clarity 3.12 PT Study Closing Date: Scheduled PT Study: The calendar date for by which all participating laboratories must submit analytical results for a PT sample to a PT Provider. | Persuasive | Suggested edits made.
| Affirmative with Comment | V1M1 3.15 | It is not appropriate to put a requirement in a definition. Reword to say "...where the study has the same pre-defined opening and closing dates for all participants..." | Persuasive | Suggested edits made. Scheduled Proficiency Testing Study (Scheduled PT Study): A single complete sequence of circulation and scoring of proficiency testing samples to all participants in a proficiency test program. Where the study has the same pre-defined opening and closing dates for all participants.
| Negative with Comment | V1M1 3.16 | Since what happens next after suspension is not defined, and reinstatement time frames are being defined in V2M2 6.3, no time frames are needed here. For example, even if a lab were suspended indefinitely due to failure to correct deficiencies, the effect would be the same as if accreditation were revoked. The period of suspension is better left to the governing laws, regulations and procedures of the AB. In fact, this language could imply that once the suspension period expires according to these timelines, the accreditation would be reinstated regardless of the lab’s compliance status, which clearly is not the intent. 3.16 Suspension: The temporary removal of a laboratory’s accreditation for a defined period of time, which shall not exceed six (6) months or the period of accreditation, whichever is longer, in order to allow the laboratory time to correct deficiencies or area of non-conformance with the Standard. | Persuasive | Suggested edits made. 3.16 Suspension: The temporary removal of a laboratory’s accreditation for a defined period of time.
| Affirmative with Comment | V1M1 3.16 | Section 3.16 Suspension: This definition is not all inclusive as to what may lead to a suspension of accreditation. There are other conditions for accreditation that if not met will lead to suspension. These can include lack of payment of fees, not undergoing an assessment in a timely manner, misrepresenting their accreditation in reporting or marketing materials etc. In fact, it implies that if a lab has deficiencies and needs to initiate corrective action the lab will be suspended, which is not accurate. | Persuasive | Edits made; definition will now be more generic as details provided within the standard itself.
| Affirmative with Comment | V1M1 3.16 | The way that "Suspension" is defined, on the "whichever is longer," can a laboratory in fact stay suspended for 6-12 months if its accreditation expires after the 6 months when a suspension would otherwise start? In other words, if the accreditation period is January 1 – December 31, and the laboratory fails its second PT (on latest 3 attempts) on January 2, can the laboratory in fact remain suspended ALL THE WAY to December 31? If this not the intent, the recommended correction: 3.16 Suspension: … period of accreditation, whichever is shorter, in order to allow the …" | Persuasive | New language to read Suspension: The temporary removal of a laboratory’s accreditation for a defined period of time, which shall not exceed six (6) months or the period of accreditation, whichever is shorter, in order to allow the laboratory time to correct deficiencies or area of non-conformance with the Standard.
| Affirmative with Comment | V1M1 3.2 | Does not match the definition in V2M2. These should be consistent. | Persuasive | Definition removed
| Affirmative with Comment | V1M1 3.6 | Makes no sense...what is an "unknown sample"? revise to use ISO/IEC 17043 definition of PT | Persuasive | While the committee cannot use language from ISO/IEC 17043, the definition is reworded for clarification. “Proficiency Testing (PT): A means to evaluate a laboratory’s performance under controlled conditions relative to a given set of criteria, through analysis of samples whose results are unknown to the laboratory and are provided by an external source.”
| Affirmative with Comment | V1M1 3.8 | Definition 3.8 and requirement 4.1.3 are redundant. The definition 3.8 should be revised to state what a PT Provider is, the requirement to be accredited is covered in 4.1.3 | Persuasive | Changed language to: A person or organization which operates a PT program.
| Affirmative with Comment | V1M1 3.10 | 3.10 Proficiency Testing Reporting Limits (PTRL)…The PTRL is defined in terms of the minimum concentration of an analyte prepared by the PTP. I suggest adding a sentence that explains the PTRL from the lab’s perspective, e.g., “The PTRL is also the minimum concentration to which laboratory PT results must be reported.” | Non-Persuasive | The PTRL is explained in length in sections 4.2 and 4.3 (4.2.3, 4.3.5). 4.3.7 states If the value is < LOQ and the LOQ is less than the PTRL, the laboratory may choose to evaluate results to the LOQ rather than the PTRL. If the value found is
| Affirmative with Comment | V1M1 3.10 | Remove word "theoretical" ... the concept "theoretical concentration" is nonsense | Persuasive | Removed the word theoretical from the definition.
| Affirmative with Comment | V2M2 3.8 | Definition 3.8 and requirement 4.1.3 are redundant. The definition 3.8 should be revised to state what a PT Provider is, the requirement to be accredited is covered in 4.1.3 | Persuasive | Changed language to: A person or organization which operates a PT program.
| Affirmative with Comment | V1M1 3.10 | Remove word "theoretical" ... the concept "theoretical concentration" is nonsense | Persuasive | Removed the word theoretical from the definition.
| Negative with Comment | V1M1 4.2 | Sections 4.2 and 4.3 indirectly mandate that all labs must report to or below the PTRL. This contradicts the requirement that labs manage samples in the same manner as routine samples and places an undue burden on laboratories. | Non-Persuasive | In order for results to be evaluated consistently, the PT Provider instructions must be followed in preparing dilutions and establishing testing conditions. The LOQ is based on the dilutions specified in the instructions, so if instructions are followed, the reporting requirement does not add any additional burden to the lab.
| Negative with Comment | V1M1 4.3 | Sections 4.2 and 4.3 indirectly mandate that all labs must report to or below the PTRL. This contradicts the requirement that labs manage samples in the same manner as routine samples and places an undue burden on laboratories. | Persuasive | The section is now 4.2.4 and has been reworded to allow laboratories to report results to their LOQ.
| Affirmative with Comment | V1M1 4.4 | This entire section duplicates section 4.3.6. Suggest removing either 4.4 or 4.3.6. | Persuasive | Agreed that it's redundant to include this statement in two different sections. And as a record-keeping requirement, this is a better fit in 4.4 (not in 4.3 for reporting requirements). Removing section 4.3.6. Note: also changed numbering of 4.3.7 to 4.3.6 and 4.3.8 to 4.3.7
| Negative with Comment | V1M1 6.1 | I suggest that this section should be separated into sub-sections for readability and corrected for some grammar. Such as: 6.1 A laboratory that fails to successfully analyze a PT study for a particular FOA shall determine the root cause of the failure and take corrective action. 6.2 The laboratory shall document the root cause investigation and subsequent corrective action. NOTE: The requirements for corrective action are described in V1M2 of this standard. 6.3 The laboratory shall provide the root cause investigation and corrective action documentation to the AB within 30 calendar days of a request from the AB. 6.4 Failure to submit documentation of the root cause investigation or corrective action, or both, to the AB within 30 calendar days of the request from the AB is due cause for suspension of accreditation for a particular FOA. | Persuasive | Reformated.
| Negative with Comment | V1M1 6.1 | An AB may require labs to always submit these records. 6.1 If the laboratory fails to successfully analyze a PT study for a particular field of accreditation, it shall determine the root cause for the failure and take any necessary corrective action. The laboratory shall document the investigation and corrective action. The requirements for corrective action are described in Volume 1 Module 2 of this standard. The laboratory shall provide these records to the primary AB within thirty (30) calendar days upon receipt of a request by the AB. If the AB institutes a permanent request for these records (i.e., the laboratory must submit documention of the investigation and corrective action for every PT failure), then the laboratory shall provide them to the primary AB within thirty (30) calendar days of the investigation and corrective action. Failure to submit a corrective action records within thirty (30) calendar days of a documented request from the primary AB may result in suspension. | Non-Persuasive | We cannot add "possible" requirements to the standard. Unless all ABs institute a requirement for PT corrective action documentation, this cannot be added. However we can add the word "record" to the last statement and remove the word "a".
| Affirmative with Comment | V1M1 6.1 | “…The laboratory shall provide these records to the primary AB within thirty (30) calendar days upon receipt of a request by the AB. Failure to submit a corrective action within thirty (30) calendar days of a documented request from the primary AB may result in suspension. “ Comment: Why is it “upon receipt of a request by the AB” in the first case, and then “…a documented request from the primary AB” later? Suggest using consistent verbiage “primary AB” throughout this section, if that was what was intended. | Persuasive | Added the word "primary".
| Affirmative with Comment | V1M1 7 | Recommend adding a Section 7.2 regarding how to submit questions/complaints about an ABs evaluation of PT data. There is a corresponding section in V2M2 (section 5.2.1) that ABs must have procedures to respond to lab questions about validity of an ABs evaluation of lab's performance on a PT Study. | Persuasive | New section added to V1M1. 7.2 - Laboratories shall have submit questions to their AB in regards to the AB's PT evaluation, if necessaary.
| Negative with Comment | V1M1 1.3.2 | Section 1.3.2 is an incomplete. It must include the following statement. "For Fields of Accreditation (FOA) that are not designated as FoPTs by this Standard, the applicable regulation, method or contract requirements will apply." | Non-Persuasive | Out of scope of the standard.
| Negative with Comment | V1M1 1.3.2 | 1.3.2 reads as follows. This Standard applies only to Fields of Accreditation (FOA) that are designated as FoPT by the TNI PTPEC. It should read as follows. This Standard applies only to FOA that are designated as FoPT by the TNI PTPEC. For FOA that are not designated as FoPTs, the appropriate protocol, method, regulation or contract applies. Where there are no such requirements in a method, regulation or contract, the laboratory shall document in its QA Manual it is own criteria for analyzing external PT samples. | Non-Persuasive | Out of scope of the standard.
| Affirmative with Comment | V1M1 1.3.2 | The PTPEC only designates Fields of Proficiency Testing (FoPTs). It does NOT designate which Fields of Accreditation (FoAs) require participation in a specific FoPT. That is the role of the environmental laboratory's Primary NELAP Accreditation Body. Recommended language, after change: 1.3.2 This standard applies only to fields of accreditation (FOA) that are associated with fields of proficiency testing (FoPT) designated by the TNI Proficiency Testing Program Executive Committee (PTPEC). | Persuasive | Updated as suggested. 1.3.2 This standard applies only to fields of accreditation (FOA) that are associated with fields of proficiency testing (FoPT) designated by the TNI Proficiency Testing Program Executive Committee (PTPEC).
| Negative with Comment | V1M1 3, 4.2.3, 4.3.5 and 4.3.7 | The term LOQ should be included in the Terms and Definitions section. Sections 4.2.3, 4.3.5 and 4.3.7 a) through c) refer to the LOQ. The LOQ is not typically determined for radiochemical analyses, thus comparisons to the LOQ cannot be made. Section 4.3.7 c) states that results shall be censored if the result is less than the PRRL or the LOQ (reported as "<" the PTRL or LOQ). Section 1.7.2.4 b) of Module 6 (EL-V1Mg-2012 Rev0.1) states that for radiochemical measurements "Each result shall be reported with its measurement uncertainty." This statement is reiterated in Section 1.5.4 a) of the proposed draft of module 6. There are no provisions for censoring radioanalytical data. | Persuasive | All references to LOQ for radiochemistry are removed throughout the document. Added section after 4.3.7 stating : For radiochemistry, all results, including negative numbers, shall be reported whether above or below the MDA. Each result shall be reported with its Combined Standard Uncertainty (CSU). At no time shall radiochemistry results be reported as "<" the MDA.
| Affirmative with Comment | V1M1 3.0 | WET as the acronym for Whole Effluent Testing is not properly defined. | Non-Persuasive | Non persuasive- the acronymn definition for WET is "Whole Effluent Toxicity" as defined in EPA methods
| Affirmative with Comment | V1M1 4.0 | The heading for section 4.0 is titled "Requirements for Accreditation", which presumably contains all such requirements. However, heading for section 5.0 is then titled "PT Study Frequency Requirements for Accreditation". It seems more logical and easier to understand if all of the 'requirements for accreditation' were under one such heading (PT study frequency requirements were simply listed as another sub-heading under 4.0). Or alternatively change the 5.0 heading to "PT Study Frequency". | Persuasive | Changed the header to read PT Study Frequency.
| Affirmative with Comment | V1M1 4.1.2 | The sentence is really confusing. Proficiency testing is required for FoPTs listed in the TNI FoPT Tables. These Tables inherently have NO relation to Fields of Accreditation until NELAP Recognized ABs specify their Fields of Accreditation available for accreditation and then what the FoPTs are required (if at all) for FOAs. Recommended language after changes are made: 4.1.2 The laboratory shall participate in PT studies for each field of accreditation where corresponding FoPTs exist in the TNI FoPT tables, for which the laboratory seeks to obtain or maintain accreditation. | Persuasive | Changed 4.1.2 to match other edits.
| Negative with Comment | V1M1 4.1.2 | Clarification 4.1.2 The laboratory shall participate in PT studies for each field of accreditation where corresponding FoPTs exist in the TNI FoPT tables and for which the laboratory seeks to obtain or maintain accreditation. | Persuasive | Changed 4.1.2 to match other edits.
| Affirmative with Comment | V1M1 4.1.3 | For clarity, the comma in the sentence should be deleted and an extra word should be inserted as shown below. Recommended appearance after the change is made: 4.1.3 The laboratory shall obtain scheduled PT studies or supplemental studies for the individual fields of proficiency testing from a PT Provider accredited by a TNI approved PTPA. | Persuasive | Changed 4.1.3 to read - "4.1.3 The laboratory shall obtain scheduled PT studies or supplemental studies for the individual fields of proficiency testing from a PT Provider accredited to Volume 3 of this Standard by a TNI approved PTPA
| Affirmative with Comment | V1M1 4.1.3 | This clause as written does not require that the PT Provider be TNI-accredited, only the PTPA. There are PT Providers that are accredited by A2LA and ACLASS (which are TNI approved PTPAs), but that are not TNI PT Providers. Proposed update: "…from a PT Provider accredited to Volume 3 of this Standard by a TNI approved PTPA." | Persuasive | Changed 4.1.3 to read - "4.1.3 The laboratory shall obtain scheduled PT studies or supplemental studies for the individual fields of proficiency testing from a PT Provider accredited to Volume 3 of this Standard by a TNI approved PTPA
| Affirmative with Comment | V1M1 4.1.3 | The definition of “PT Study” in the Standard “refers to a Scheduled PT Study or a Supplemental PT Study.” Therefore the language in 4.1.3 that refers to “PT studies or supplemental studies” is redundant. Suggest removing “or supplemental studies.” | Persuasive | Removed "or supplemental"
| Affirmative with Comment | V1M1 4.1.4 | The sentence has an incorrect section reference. The sentence should also make mention of the required frequency of participation in PT studies. The NOTE in the section also need improvement. Recommended appearance after corrections is made: 4.1.4 T | Persuasive | Edits as suggested were added.
| Affirmative with Comment | V1M1 4.1.4 | The FoPT tables themselves do not require PT studies. The FoPT tables merely contain the “matrix, technology/method, analyte combinations for which the composition, spike concentration ranges and acceptance criteria have been established by the PTPEC,” t | Persuasive | Language edited
| Affirmative with Comment | V1M1 4.1.4 | 40 CFR 141 is for "regulated" EPA analytes, not necessarily all potable water PTs. There are PW analytes/PTs that are not part of 40 CFR 141 and would therefore not require PTs by test method. There could also be additional PW anatlyes added to the FoPT table in the future that are not part of 40 CFR 141. Proposed update: "Note: PT results are required by Federal Regulations, 40 CFR 141, per test method rather than technology for applicable potable water PTs." | Persuasive | Language edited
| Affirmative with Comment | V1M1 4.1.5 | The FoPT tables themselves do not require PT studies. The FoPT tables merely contain the “matrix, technology/method, analyte combinations for which the composition, spike concentration ranges and acceptance criteria have been established by the PTPEC,” to quote the definition in the Standard. Suggest changing 4.1.4 to: “The laboratory shall analyze unique, single blind, single concentration PT samples, when defined in the TNI FoPT tables described in section 4.1.1, to determine compliance…” | Non-Persuasive | The requirment addresses any personell within the laboratory and across a laboratory network.
| Affirmative with Comment | V1M1 4.2.3 | The sentence is long and might be better broken into two shorter sentences. Also, the acronym LOQ does not appear to be defined. Recommended wording after change is made: 4.2.3 “… established by the TNI FoPT Tables. Alternatively, if the laboratory's LOQ (Limit of Quantitation) is below ...” | Persuasive | Dividing 4.2.3 into two sentences- wording flows better. LOQ may be defined in the main glossary. Will verify, or add if it is not included in the glossary.
| Negative with Comment | V1M1 4.2.3 | Define LOQ acronym. 4.2.3 The laboratory shall evaluate the analytical result for each chemistry and radiochemistry field of accreditation to the PTRL as established by the TNI FoPT Tables, or if the laboratory’s Limit of Quantitation (LOQ) is below the PTRL, they may evaluate results to their normal LOQ. | Persuasive | If LOQ is not writtten completely previously in the standard this comment is persuasive, and the proposed language will be added.
| Negative with Comment | V1M1 4.2.3, 4.3.5 and 4.3.7 | Negative on Sections 4.2.3, 4.3.5 and 4.3.7a-c, including the subsequent **Note. Rationale: The practices proposed are inconsistent with the protocols and requirements for reporting in Module 6 (EL-V1M6) of the TNI Standard (Radiochemistry). Radiochemical results should only be reported “as measured” and in association with the measurement uncertainty. Censoring results, reporting results without uncertainty, or classifying results as unacceptable based on their concentration fails to recognize the statistical nature of radiochemical measurement results. • This requirement requires laboratories to censor results below the LOQ. o The LOQ concept is not defined for radiochemistry. o Detection capability concepts such as the Critical Value, the MDC, and the SDWA Detection Limit, are defined (see proposed working draft standard - EL-V1M6 WDS-Rev0-5-29-14) but these should not be used to censor results. • Radiochemical measurement results must be reported in association with their measurement uncertainty. o Reporting radiochemical results in association with their uncertainty ensures that they will be valid in spite of the concentration measured in samples; o Both the currently approved version and proposed working draft of the TNI Standard Radiochemistry Module (Module 6 - EL-V1M6) stipulate that radiochemical results be reported in association with the associated measurement uncertainty and that the type of uncertainty (counting or total) and the confidence level or coverage factor for the uncertainty be specified (ensures the results can be interpreted unambiguously); o The proposed working draft standard of Module 6 (WDS-Rev0-5-29-14) also clearly stipulates that results should be reported even if they are negative. In other words, no result should be censored based on the amount of activity measured. | Persuasive | Language re-written
| Negative with Comment | V1M1 4.2.3, 4.3.5 and 4.3.7 | The Voting draft is incomplete from the Radiochemical Testing perspective. Lumping radiochemical testing with chemical testing into one under Proficiency Testing is problematic. Of particular concern are the sections 4.2.3, 4.2.4, 4.3.5 and 4.3.7 a-c and **Note. These are inconsistent with the requirements and practices of Radiochemical Testing. Terms such as PTRL, LOQ are not applicable for Radiochemical measurements, so also the term re-scaling initial calibration curve. Additionally, Safe Drinking Water Act (SDWA) Detection Limits is completely omitted in the Standard. In the present form, the Standard is unacceptable. | Persuasive | Language re-written
| Negative with Comment | V1M1 4.2.3, 4.3.5 and 4.3.7 | Sections 4.2.3, 4.3.5 and 4.3.7 a) through c) refer to the LOQ. The LOQ is not typically determined for radiochemical analyses, thus comparisons to the LOQ cannot be made. | Persuasive | Language re-written
| Negative with Comment | V1M1 4.2.4 | These sections refer to calibration curves defined as a function of concentration which does not applicable to radiochemistry. | Persuasive | Removed the word radiochemistry from section 4.2.4.
| Negative with Comment | V1M1 4.2.4a | This practice encourages laboratories to deviate from their standard operating procedure. This would not be done on a normal environmental sample and would thus not be a good indicator or how normal environmental samples are analyzed. | Non-Persuasive | SOP's should have the flexablity to allow for samples that fall outside their "normal" or expected concentration range. A good SOP should reflect how this is accomplish this.
| Negative with Comment | V1M1 4.2.4a | This practice encourages laboratories to deviate from their standard operating procedure. This would not be done on a normal environmental sample and would thus not be a good indicator or how normal environmental samples are analyzed. | Non-Persuasive | SOP's should have the flexablity to allow for samples that fall outside their "normal" or expected concentration range. A good SOP should reflect how this is accomplish this.
| Negative with Comment | V1M1 4.2.4b | Results reported outside of the calibration range can only be estimates and thus are not a good indicator of accuracy of results from environmental samples not reported at this level. | Persuasive | The cross reference is changed from 4.3.5 to 4.3.7.
| Affirmative with Comment | V1M1 4.2.4b | There is a cross reference to section 4.3.5 with respect to reporting PT results below the calibration range. It appears that the more appropriate cross reference is section 4.3.7, which discusses FoPT reporting. | Persuasive | The cross reference is changed from 4.3.5 to 4.3.7.
| Negative with Comment | V1M1 4.3.2 | Need to allow this on the closing date. Need to specify Primary AB, since it is the only one that really matters. 4.3.2 The laboratory shall, on or before the closing date of the study, direct the PT Provider to report the PT study performance results directly to the Primary and other selected AB(s). For initial accreditation(s), the laboratory shall direct the PT Provider to provide all relevant PT Study results to the AB to support their its accreditation application. | Persuasive (first comment) Non-Persuasive (second comment) | 1st comment: Need to allow labs to select ABs on the closing date as well as before the closing date. Changed wording to suggested. 2nd comment: It doesn't matter what ABs are selected and is completely up to the lab. Expanded to say “…directly to the AB(s) designated by the laboratory…” Changed "their" to "its" as suggested.
| Affirmative with Comment | V1M1 4.3.2 | “…The laboratory shall, before the closing date of the study, direct the PT Provider to report the PT study performance results directly to the selected AB(s)…” Comment: What are “selected AB(s)?” Suggest expanding this to say “…directly to the AB(s) designated by the laboratory…” | Persuasive | changed wording to suggested.
| Affirmative with Comment | V1M1 4.3.3 | The exact meaning and intent of this sentence is unclear to me. I would guess that the phrase “and criterion that identifies the laboratory for the Field of Accreditation for which the PT sample was analyzed” refers to method codes and similar designations used by the ABs, but that isn’t obvious. Please clarify. | Persuasive | The language was changed to clarify the reporting requirements for the laboratory.
| Affirmative with Comment | V1M1 4.3.4 | Again, as in the comment above on section 4.1.4, the reference to drinking water is mis-leading. It makes it sound like all drinking water analytes must be reported per test method, when it's really only the regulated analytes in 40 CFR 141 that must be reported by test method. Proposed update: Except for drinking water analytes referenced in 40 CFR 141, a laboratory may choose..." | Persuasive | changed wording to suggested.
| Negative with Comment | V1M1 4.3.5 | 4.3.5 Limits of quantitation should be established by what is required to produce high quality analytical results at the levels needed by the data user. The practice of TNI establishing limits encourages labs to deviate from their standard operating procedures and treat samples other than the way they treat normal environmental samples. If a state or regulatory body would like to set limits this should be covered under section 4.1.7 and not a rule implemented across the board. | Non-Persuasive | In order for results to be evaluated consistently, the PT Provider instructions must be followed in preparing dilutions and establishing testing conditions. The LOQ is based on the dilutions specified in the instructions, so if instructions are followed, the reporting requirement must be followed. In the case of PT samples, the PT Provider is the "client" and determined their reporting requirements, as any regualtory program or client would do.
| Negative with Comment | V1M1 4.3.5 | 4.3.5 Limits of quantitation should be established by what is required to produce high quality analytical results at the levels needed by the data user. The practice of TNI establishing limits encourages labs to deviate from their standard operating procedures and treat samples other than the way they treat normal environmental samples. If a state or regulatory body would like to set limits this should be covered under section 4.1.7 and not a rule implemented across the board. | Non-Persuasive | In order for results to be evaluated consistently, the PT Provider instructions must be followed in preparing dilutions and establishing testing conditions. The LOQ is based on the dilutions specified in the instructions, so if instructions are followed, the reporting requirement must be followed. In the case of PT samples, the PT Provider is the "client" and determined their reporting requirements, as any regualtory program or client would do.
| Affirmative with Comment | V1M1 4.3.6 | The acronym "AB" shows up by itself and is not associated with "Primary AB" as defined earlier. Add the acronym to the definition of "Accreditation Body" in clause 3.1: 3.1 "Accreditation Body (AB): The territorial, state or federal …" | Persuasive | Added acronym (AB) to definition in 3.1 Doing the same for V2M2 definition for consistency. Reviewed both volumes to ensure that "Primary AB" and "AB" are used as intended since they have differenct definitions.
| Negative with Comment | V1M1 4.3.6 | This is a record retention requirement and should be included in section 4.4. But, a very similar statement is already part of section 4.4.1. | Persuasive | Identical statement in 4.4.1; removed 4.3.6 under reporting requirements.
| Negative with Comment | V1M1 4.3.7 | I think this entire section is very confusing and actually contains information about how PTs will be scored depending on how they are reported rather than actual standard requirements for PTs. The committee could consider just telling the laboratory to report values at or below the PTRL as established in the TNI FoPT Tables. Alternatively, stating: “A laboratory that reports a value with a less than (<) symbol that is higher than the PTRL where the acceptance range of the PT study is a measurable value will result in a “Not Acceptable” score. “ | Persuasive | 4.3.7 The laboratory shall evaluate and report each FoPT result as follows: a) If the value found by the laboratory is equal to or above the PTRL and equal to or above the LOQ, the laboratory shall report the value found as the analytical result. b) If the value found by the laboratory is equal to or above the PTRL and less than the LOQ, the laboratory shall report the result without the qualification of result required in Volume 1 Module 4 of this Standard or the laboratory shall choose to evaluate results to the LOQ rather than the PTRL.** c) If the value found is less than the PTRL, the laboratory shall report a result of “<” the PTRL value, a result between the LOQ (if below PTRL) and PTRL, or a result less than (<) laboratory’s LOQ. The PTRL value shall not be adjusted for sample amount used or percent moisture. ** Note: In the case where the laboratory LOQ is greater than the PTRL: If the laboratory chooses to report a value of < LOQ and the analyte is present above the PTRL, the result will be scored as “Not Acceptable” by the PT provider.
| Affirmative with Comment | V1M1 4.3.7a | "V1M4" shows up in the sentence. Does everyone in TNI know what this means? Spell out V1M4 as "Volume 1, Module 4" of this Standard. | Persuasive | The requested change was made to the standard.
| Negative with Comment | V1M1 4.3.7a | Results reported outside of the calibration range can only be estimates and thus are not a good indicator of accuracy of results from environmental samples not reported at this level. | Non-Persuasive | The Committee had extensive discussions concerning this issue and decided to provide flexibility in 4.3.7 (including the note in 4.3.7) as an option to any laboratory that did not want to change their normal procedures to report to the PTRL.
| Negative with Comment | V1M1 4.3.7a | Results reported outside of the calibration range can only be estimates and thus are not a good indicator of accuracy of results from environmental samples not reported at this level. | Non-Persuasive | The Committee had extensive discussions concerning this issue and decided to provide flexibility in 4.3.7 (including the note in 4.3.7) as an option to any laboratory that did not want to change their normal procedures to report to the PTRL.
| Negative with Comment | V1M1 4.3.7c | The note indicating that "In the case where a laboratory LOQ is greater than the PTRL: If the laboratory chooses to report a value of <LOQ and the analyte is present above the PTRL, the result will be scored as "Not Acceptable" by the PT Provider. " I disagree with this approach. This punishes the laboratory for performing the analysis and reporting the results in accordance with their policies and procedures. Also, if the laboratories had a deadline with which to implement the 2009 standard requirements why is this not also applicable to the accreditation bodies. There was ample time to determine a course of action to implement the requirements. | Non-Persuasive | The Committee had extensive discussions concerning this issue and decided to provide flexibility in 4.3.7 (including the note in 4.3.7) as an option to any laboratory that did not want to change their normal procedures to report to the PTRL.
| Negative with Comment | V1M1 4.3.7c | **NOTE:) Failure of a PT based on its concentration being below the laboratories' concentration range is not a good evaluation of their performance on everyday environmental samples in which they would not report that result. A laboratory should not be assessed at levels in which they do not analyze normal environmental samples. This practice will encourage labs to deviate from their standard operating procedures just to pass a PT. A PT is supposed to be a means to evaluate a laboratory's performance. The performance that should be evaluated is that which is generated from the standard operating procedure they use everyday for client samples, not an altered procedure used to pass a PT sample that falls below their limit of quantitation. | Non-Persuasive | The Committee had extensive discussions concerning this issue and decided to provide flexibility in 4.3.7 (including the note in 4.3.7) as an option to any laboratory that did not want to change their normal procedures to report to the PTRL.
| Negative with Comment | V1M1 4.3.7c | **NOTE:) Failure of a PT based on its concentration being below the laboratories' concentration range is not a good evaluation of their performance on everyday environmental samples in which they would not report that result. A laboratory should not be assessed at levels in which they do not analyze normal environmental samples. This practice will encourage labs to deviate from their standard operating procedures just to pass a PT. A PT is supposed to be a means to evaluate a laboratory's performance. The performance that should be evaluated is that which is generated from the standard operating procedure they use everyday for client samples, not an altered procedure used to pass a PT sample that falls below their limit of quantitation. | Non-Persuasive | The Committee had extensive discussions concerning this issue and decided to provide flexibility in 4.3.7 (including the note in 4.3.7) as an option to any laboratory that did not want to change their normal procedures to report to the PTRL.
| Negative with Comment | V1M1 4.3.8 | This sentence does not belong in a standard, but if the committee chooses to include it, it shouldn’t be in the reporting requirements section. It should maybe be included in a “NOTE.” Or change the words to say, “The laboratory shall not attempt to obtain the prepared value, PT Provider’s assigned value verification, homogeneity, or stability testing of a PT study, or a portion of a PT study, from the PT Provider prior to the time the results of the study are released by the PT study Provider.” | Persuasive | Agreed, this probably does not need to be in the standard. It's not really a requirement, so it doesn't belong in the Reporting Requirements section. And it seems unnecessary to include it elsewhere in the standard as a requirement for labs to not ask for this information prior to a study close. It could be a note, but then is simply an fyi that a lab has the right to request this information. Removed section 4.3.8 from the standard.
| Affirmative with Comment | V1M1 4.3.8 | “…for which results were submitted to a laboratory AB.” Comment: Why is it “laboratory AB” here, and not just “AB” as it is elsewhere in the Standard? | Persuasive | For consistency, it makes sense to remove "laboratory". However, the section is removed in respionse to another comment.
| Affirmative with Comment | V1M1 5.1.1 | The sections addressing the laboratories ability to not run a 3rd study if the first two are passing is a little misleading. I believe a disclaimer should be added that if the next study fails, the laboratory would lose accreditation. Yes, this gives the labs opportunity to save money, but the risk seems high. | Non-Persuasive | The note provides the necessary clarity to section a
| Affirmative with Comment | V1M1 5.1.1 | “d) The closing date of the most recent successful PT study for an FoPT must be no more than six (6) months prior to the application for initial accreditation and the laboratory shall continue to participate in PT studies semi-annually from that point on. “ Comment: The last part of the requirement, “…and the laboratory shall continue to participate in PT studies semi-annually from that point on.” Is not a requirement of Initial Accreditation, it is a requirement of Continued Accreditation (as stated in 5.2.1.2). It therefore does not belong in section 5.1.1. Suggest removing that portion of the statement. | Persuasive | Requirement changed to be more clear.
| Affirmative with Comment | V1M1 5.1.1a | The reference to Section 4.1.1 appears to be incorrect. Change the section reference to "4.1.2". | Persuasive | Reference changed.
| Affirmative with Comment | V1M1 5.1.1d | The PT frequency of "semi-annually" by itself may not be enforceable by accreditation bodies on laboratories. Also, the sentence appears to penalize laboratories that, for whatever reason, may choose to do PTs more often. Proposed language to be used: “… shall continue to participate in PT studies at least semiannually (no more than 7 months apart between consecutive attempts) from that point on." | Persuasive | Language was changed to be more clear.
| Negative with Comment | V1M1 5.1.1d | Avoid varying interpretations of “semi-annually.” Refer to the frequency requirements in 5.2. 5.1.1 d) The closing date of the most recent successful PT study for an FoPT must be no more than six (6) months prior to the application for initial accreditation and the laboratory shall continue to participate in PT studies semi-annuallyat least twice per year in accordance with the frequencies given in section 5.2 of this module from that point on. | Persuasive | Language was changed to be more clear.
| Negative with Comment | V1M1 5.1.2 | This sentence should be re-worded for grammar, “For Whole Effluent Toxicity (WET) testing, the laboratory shall demonstrate to the primary AB that it has received an acceptable evaluation for at least one (1) PT study to obtain initial accreditation.” | Persuasive | Language changed
| Affirmative with Comment | V1M1 5.2.1.1 | The reference to Section 4.1.1 appears to be incorrect. Change the section reference to "4.1.2". | Persuasive | Reference changed
| Negative with Comment | V1M1 5.2.1.2 | The committee has included the term “semi-annually” in this section but has not defined this term. One definition of “semi-annually” is “occurring or issued twice a year.” The problem with using this term is that the PT frequency described in the sections that follow this term say that a laboratory must participate in PT studies that are no longer than 7 months from closing date to closing date of each PT and if the lab fails to do so, a failure will be recorded in their PT history. For example, Laboratory ABC currently has a PT history of 3 passes out of the most recent 3 attempts. Their next rounds of PTs are as follows: • Study #1 closes January 1, 2013, and the lab obtains “Acceptable” scores for all required FoPTs (new history of Pass, Pass, Pass; so 3 passes out of 3 attempts) • Study #2 is due by August 1, 2013, but the lab does not perform the study. The AB records a failure in the Laboratory ABC’s history for failure to meet the PT requirements. (new history of Pass, Pass, Fail; so 2 passes out of 3 attempts) • The next PT study is due in 7 months, or March 1, 2014 and the laboratory obtains “Acceptable” scores for all required FoPTs (new history of Pass, Fail, Pass; so 2 passes out of 3 attempts) The example above shows that a laboratory can meet the 2 Passes out of 3 Attempts without actually meeting the semi-annual schedule of PT performance. This can also result in the laboratory waiting 14 months between analyses of PT studies. The addition of “semi-annual” could allow an AB to suspend the laboratory’s accreditation for failure to analyze two PTs per year while the laboratory actually continues to meet the requirement of 2 passes out of 3 attempts. (NOTE: I completely understand that this is not really any different than the current PT requirement from the 2009 standard, but I just wanted to point out that the standard isn’t consistent as written.) | Persuasive | Language was changed to be more clear.
| Negative with Comment | V1M1 5.2.1.2 | Avoid varying interpretations of “semi-annually.” 5.2.1.2 The laboratory shall analyze and report a PT study semi-annually at least twice per year for each accreditation FoPT for which it seeks to maintain accreditation that meetsin accordance with the following criteria: | Persuasive | Language was changed to be more clear.
| Affirmative with Comment | V1M1 5.2.2 | The acronym "WET" shows up without a clear definition as to what it stands for. Define WET upon first usage, as follows: 5.2.2 For Whole Effluent Toxicity (WET) Testing: To maintain accreditation…" | Persuasive | The acronymn definition for WET is "Whole Effluent Toxicity" as defined in EPA methods. The acronymn has been defined.
| Negative with Comment | V1M1 5.2.2 | I think that laboratories should be required to pass the WETT PT. I ask that the committee reconsider the requirement to only participate. | Non-Persuasive | WET Testing requirements follow the protocols for DMR-QA, which requires corrective action. 6.2 specifies documentation required for WET failure corrective action.
| Negative with Comment | V1M1 5.2.2 | Define WET acronym 5.2.2 For Whole Effluent Toxicity (WET) Testing | Persuasive | The acronymn definition for WET is "Whole Effluent Toxicity" as defined in EPA methods. The acronymn has been defined
| Affirmative with Comment | V1M1 5.2.2a | If the laboratory participates in WET PT testing through DMRQA, is the laboratory allowed to hang on to the PT for the entire 90 days of the study? No change is necessary if Volume 3 requirements on PT study duration have not been changed, in which case the laboratory can only have the DMRQA WET PT for 45 days maximum before a result is reported to the PT Provider. | Non-Persuasive | The requirements of the DMR-QA study must be followed, and the sample can be kept for the duration the study is open.
| Negative with Comment | V1M1 8.0 | By having these sections written the way they are, they imply that the laboratory is the organization that does the “reinstating”. Maybe the committee could use the following language, “A laboratory that seeks reinstatement of accreditation after suspension due to unsatisfactory PT performance shall meet the requirements for continued accreditation as described in Section 5.2.” And, “A laboratory that seeks reinstatement of accreditation after revocation of accreditation due to unsatisfactory PT performance shall meet the requirements for initial accreditation as described in Section 5.1.” Also, this section does not include information regarding how a lab gets accreditation back for failure to submit a corrective action report as specified in section 6. | Persuasive | The proposed language better states the intent of sections 8.1 and 8.2. Also, it makes sense to address corrective action. A section 8.3 is now added in the same format as 8.1 and 8.2.
| Affirmative with Comment | V2M2 3.0 | Include date of the relevant standards (i.e. ISO/IEC 17011:2004) | Persuasive | Updated sec 3.0 ISO/IEC 17011:2004 and ISO/IEC 17025:2005
| Affirmative with Comment | V2M2 3.1 | 3.1 defines accreditation body as "territorial, state or federal agency"... if left as is, this definition will need to be revisited if and when non-governmental ABs are recognized under this program for laboratory accreditation. | Persuasive | Definition changed to be more aligned with ISO 17011.
| Affirmative with Comment | V2M2 3.1 | The definition may need to be revisited in the future if TNI proceeds with recognizing non-governmental accreditation bodies as okay for the NELAP program. Alternatively, are there plans for a different Volume 2, Module 2 for a non-state environmental laboratory TNI accreditation system? | Persuasive | Definition changed to be more aligned with ISO 17011.
| Negative with Comment | V2M2 3.1 | This definition potentially excludes Non-Governmental ABs. There is a major effort underway in TNI to recognize NGABs to accredit to the TNI Standard. 3.1 Accreditation Body: The territorial, state or federal agencyorganization having responsibility and accountability for environmental laboratory accreditation and which grants accreditation under this program. | Persuasive | Definition changed to be more aligned with ISO 17011.
| Affirmative with Comment | V2M2 3.1 | Does not match the definition in V1M1. These should be consistent | Persuasive | Definition changed to be more aligned with ISO 17011.
| Affirmative with Comment | V2M2 3.11 | The bullet for "Scheduled PT Study" should have an additional word inserted for clarification. Clause 3.11 (re-numbered) should read as follows, for the "Scheduled PT Study" bullet: “Scheduled PT Study: The calendar date for which all participating laboratories must submit analytical results for a PT sample to a PT Provider." | Persuasive | The word "participating" has beeen added.
| Negative with Comment | V2M2 3.11 | Definition of PT Study Closing Date, a) Scheduled PT Study. I suggest that you add the word “participating” to the definition. Such as, “The calendar date for which all participating laboratories must submit…” | Persuasive | The word "participating" has beeen added.
| Negative with Comment | V2M2 3.11 | Editorial and to add clarity Scheduled PT Study: The calendar date for by which all participating laboratories must submit analytical results for a PT sample to a PT Provider. | Non-Persuasive | Non-persuasive. "by" refers to the calander date; "for" refers to the subject. We believe we are addressing the subject, all participating labororatories.
| Affirmative with Comment | V2M2 3.15 | It is not appropriate to put a requirement in a definition. Reword to say "...where the study has the same pre-defined opening and closing dates for all participants..." | Persuasive | The language in 3.15 is changed to say "Scheduled Proficiency Testing Study (Scheduled PT Study): A single complete sequence of circulation and scoring of proficiency testing samples to all participants in a proficiency test program where the study must have has the same pre-defined opening and closing dates for all participants.
| Affirmative with Comment | V2M2 3.16 | Section 3.16 Suspension: This definition is not all inclusive as to what may lead to a suspension of accreditation. There are other conditions for accreditation that if not met will lead to suspension. These can include lack of payment of fees, not undergoing an assessment in a timely manner, misrepresenting the accreditation on reports or marketing, etc. In fact, it implies that if a lab has deficiencies and needs to initiate corrective action the lab will be suspended, which is not accurate. | Non-Persuasive | The definition makes no refererence to what may cause a suspension. It just says it is a temporary removal of the accreditation for a finite period of time. A deficiency may be lack of payment, or misrepresentation or the like. It does not preclude anything that was mentioned in the comment.
| Affirmative with Comment | V2M2 3.16 | The way that "Suspension" is defined, on the "whichever is longer," can a laboratory in fact stay suspended for 6-12 months if its accreditation expires after the 6 months when a suspension would otherwise start? In other words, if the accreditation period is January 1 – December 31, and the laboratory fails its second PT (on latest 3 attempts) on January 2, can the laboratory in fact remain suspended ALL THE WAY to December 31? If this not the intent, the recommended correction: 3.16 Suspension: … period of accreditation, whichever is shorter, in order to allow the …" | Persuasive | New language to read: Suspension: The temporary removal of a laboratory’s accreditation for a defined period of time, which shall not exceed six (6) months or the period of accreditation, whichever is shorter, in order to allow the laboratory time to correct deficiencies or area of non-conformance with the Standard.
| Affirmative with Comment | V2M2 3.2 | The definition for Accreditation Field of Proficiency Testing pertains to more than just an Analyte. Revise the beginning of the definition, to read as follows: 3.2 “Accreditation Field of Proficiency Testing: The matrix-technology/method-analyte combination for which a laboratory is required to successfully …" | Persuasive | Definition removed
| Negative with Comment | V2M2 3.2 | Definition of “Accreditation FoPT”. This definition should be removed since it isn’t necessary since you define FoPT and since experimental FoPT is not longer used. | Persuasive | Definition removed
| Negative with Comment | V2M2 3.2 | This should not be analyte-specific only as implied here 3.2 Accreditation Field of Proficiency Testing: Analytes The matrix-technology/method-analyte combinations for which a laboratory is required to successfully analyze a PT samples in order to obtain or maintain accreditation, collectively defined as: matrix, technology/method, and analyte. | Persuasive | Definition removed
| Affirmative with Comment | V2M2 3.2 | Does not match the definition in V1M1. These should be consistent. | Persuasive | Definition removed
| Affirmative with Comment | V2M2 1.3.2 | The PTPEC only designates Fields of Proficiency Testing (FoPTs). It does NOT designate which Fields of Accreditation (FoAs) require participation in a specific FoPT. That is the role of the environmental laboratory's Primary NELAP Accreditation Body. Recommended language, after change: 1.3.2 This standard does not apply to fields of accreditation (FOA) that are not associated with fields of proficiency testing (FoPT) designated by the TNI Proficiency Testing Program Executive Committee (PTPEC). | Persuasive | Updated to: 1.3.2 This standard applies only to fields of accreditation (FOA) that are associated with fields of proficiency testing (FoPT) designated by the TNI Proficiency Testing Program Executive Committee (PTPEC).
| Negative with Comment | V2M2 1.3.2 | Section 1.3.2 is incomplete and problematic. It should also include the following. For fields of accreditation that are not designated as FoPTs, the requirements imposed by the applicable regulation, method or contract will be applicable." | Non-Persuasive | Out of scope of the standard.
| Negative with Comment | V2M2 1.3.2 | This section should be re-worded to the affirmative, not the negative. Such as, “This standard applies to the fields of accreditation that are designated as FoPTs by the TNI PT Program Executive Committee.” | Persuasive | Updated to: 1.3.2 This standard applies only to fields of accreditation (FOA) that are associated with fields of proficiency testing (FoPT) designated by the TNI Proficiency Testing Program Executive Committee (PTPEC).
| Negative with Comment | V2M2 4.2.1 | 4.2.1-3: The intended outcome is unclear. | Persuasive | The Secondary AB shall have procedures in place to communicate with the Primary AB and update a laboratory’s accreditation status based on the accreditation granted by a Primary AB.
| Negative with Comment | V2M2 4.2.2 | The Secondary AB may choose to evaluate secondary accredited and applicant laboratories’ PT study results and grant, deny, suspend, or revoke a laboratory’s accreditation as if it were the Primary AB. Comment if you say they should accept to primary evaluation in 4.2.1 and they can't add any additional requirements what is the purpose of this? Inconsistency? If all primary ABs evaluate based on standards why would this be allowed at all? | Non-Persuasive | The issue is related to how quickly Primary ABs evaluate the PT status of a laboratory. Some secondaries want to know faster than a primary can evaluate and may opt to look at them themselves. If they come across any issues they need to contact the Primary AB and get their agreement before they can proceed to any actions.
| Negative with Comment | V2M2 4.2.2 | Replacing the requirement that a Secondary AB shall accept the assessment decisions made by the Primary AB in relation to PTs with section 4.2.2 that allows a Secondary AB to make it's own evaluation of PT results and accreditation decisions based on PT results is entirely contrary to the fundamental concept of the Primary/Secondary AB system. If the Secondary AB can essentially override the Primary AB's PT evaluation then the process is no longer standardized and the Standard in essentially no longer a Standard. More specifically, the NELAP policy, as stated on the TNI website NELAP-Recognized Accreditation Bodies page, reads "A NELAP Accreditation Body will accept by recognition, the accreditation status of a laboratory issued by another NELAP Accreditation Body (this is called secondary accreditation). Each Accreditation Body must adopt and adhere to this principle as a condition of membership in NELAP". | Non-Persuasive | Non-persuasive - The issue is related to how quickly Primary ABs evaluate the PT status of a laboratory. Some secondaries want to know faster than a primary can evaluate and may opt to look at them themselves. If they come across any issues they need to contact the Primary AB and get their agreement before they can proceed to any actions.
| Affirmative with Comment | V2M2 4.2.2 | “…and must contact the Primary AB to confirm their evaluation prior to taking action on a laboratory.” Comment: It is unclear what “…to confirm their evaluation…” means. Does this mean that the Primary AB has an advise-and-consent role in this, or is the Secondary AB only notifying the Primary AB? Suggest changing the language to make clear what is intended to happen. | Persuasive | Replaced "confirm" with the word "verify".
| Negative with Comment | V2M2 4.2.3 | Any AB is allowed to impose additional requirements on laboratories as their own state rules and regulations allow. Including this statement in the standard does not negate this authority. I suggest that it does not add value and should be removed. | Persuasive | Language modified to read “The Secondary AB shall not impose additional proficiency testing requirements for FOPT’s covered by the Standard as a requisite for initial or continued NELAP accreditation.”
| Affirmative with Comment | V2M2 4.2.3 | “The Secondary AB shall not impose additional requirements for proficiency testing that are not included in this Standard as a requisite for initial or continued accreditation. “ Comment: VDS V1M1 4.1.7 states “When a regulatory program requires more stringent requirements than the requirements of this module, the laboratory shall follow the more stringent requirements.” That suggests that the AB (who may operate their own “regulatory program”) may require more stringent requirements than are stated in V1M1. However, V2M2 4.2.3 says that the Secondary AB shall not impose additional requirements that are not in this Standard. If the Primary AB did impose more stringent requirements, as allowed by V1M1 4.1.7, but the Secondary AB cannot do the same, under V2M2 4.2.3, then the situation is established whereby the Primary AB can hold laboratories to higher requirements than the Secondary AB. | Non-Persuasive | V1M1 is written for laboratories analyzing PT samples. Section 4.1.7 was written in recognition of the fact that laboratories may hold certification with any number of regulatory authorities, both within, and outside of TNI. Should any of the "non-TNI" regulatory authorities have more stringent requirements than those listed in V2M2, the more stringent requirements must be met for PTs analyzed for those authorities. V1M1 4.1.7 If a regulatory program has non-TNI FoPT proficiency testing requirements the laboratory shall comply with those requirements.
| Negative with Comment | V2M2 5.1.3 | I do not agree with section 5.1.3 5.1.3 If the AB discovers that the laboratory analyzed QC standards that a PT Provider suggested or directed a laboratory to purchase with the PT study samples for the purpose of improving the laboratory’s performance for a given PT sample It seems that this eliminates QC samples even if it is a normal part of the laboratory's procedure. It reads different than the section 6.2 that was removed. or that the PT provider has given the laboratory analysis instructions beyond those specified in this Standard, the AB shall report the results of their findings to the PT Provider’s PTPA. There are no specified instructions in this standard as to what is correct for PT providers to supply. Given this wording, even preparation instructions for the sample would violate this section. | Non-Persuasive | Non-Persuasive: Section 5.1.3 requires the AB to report to the PT Provider's PTPA only in the event that QC standards are analyzed for the specific purpose of improving the laboratory's performance for PT samples and the laboratory was directed to purchase the QC standards for this purpose by the PT Provider. If it is a laboratory's routine practice to include QC samples with each analytical batch, these QC samples may be included with analytical batches containing PT samples and would not be subject to reporting under 5.1.3. PT providers must be compliant with ISO 17043. ISO 17043 includes language specifying the instructions that must be provided to study participants. Preparation instructions are included in section 4.6.1.2c of ISO 17043 and therefore would not be a violation of section 5.1.3.
| Affirmative with Comment | V2M2 6.2.1b | There are incorrect references to sections in this sentence. Revise the sentence to read: "A laboratory violates the previsions outlined in Volume 1 Module 1 Sections 4.1.2 – 4.1.7 of this Standard." | Persuasive | There is not a section 4.1.8; changed to 4.1.7.
| Affirmative with Comment | V2M2 6.2.1b | Needs to be revised to reference Sections 4.1.4-4.1.7, not 4.1.8. | Persuasive | There is not a section 4.1.8; changed to 4.1.7.
| Negative with Comment | V2M2 6.2.1b | Correction for renumbered sections 6.2.1 b) A laboratory violates the provisions outlined in Volume 1 Module 1 Sections 4.1.4 2 – 4.1.8 7 of this Standard. | Persuasive | There is not a section 4.1.8; changed to 4.1.7.
| Negative with Comment | V2M2 6.2.1b | Refers to V1M1 Sections 4.1.4-4.1.8 as grounds for revocation. | Persuasive | There is not a section 4.1.8; changed to 4.1.7.
| Affirmative with Comment | V2M2 6.2.1b | There is no section 4.1.8 in V1M1. Update Section reference. | Persuasive | There is not a section 4.1.8; changed to 4.1.7.
| Affirmative with Comment | V2M2 6.3.1 | The sentences appear to read as statements that may or may not be true in fact. If these are supposed to be requirements of the AB, then start the sentence in each clause as: "The AB shall reinstate …" (no s after reinstate) | Persuasive | Requirements need the word "shall" for clarity. The "s" needs to be removed as it is a typo.
| Negative with Comment | V2M2 6.3.1 | Correction to use “shall” 6.3.1 The AB shall reinstates accreditation of a laboratory whose accreditation is suspended, as specified in section 6.1.1a, when the laboratory establishes a history of two (2) successful PT study results out of the most recent three (3) attempts for the particular field of accreditation, as specified in Volume 1 Module 1 of this Standard. | Persuasive | Requirements need the word "shall" for clairity. The "s" needs to be removed as it is a typo.
| Affirmative with Comment | V2M2 6.3.2 | The sentences appear to read as statements that may or may not be true in fact. If these are supposed to be requirements of the AB, then start the sentence in each clause as: "The AB shall reinstate …" (no s after reinstate) | Persuasive | Requirements need the word "shall" for clairity. The "s" needs to be removed as it is a typo.
| Negative with Comment | V2M2 6.3.2 | Correction to use “shall” 6.3.2 The AB shall reinstates accreditation of an environmental laboratory whose accreditation was suspended as specified in section 6.1.1b, when the laboratory provides the corrective action report for the failed PT study and the Primary AB determines that corrective action acceptable, provided the laboratory meets the requirements for continued accreditation. | Persuasive | Requirements need the word "shall" for clairity. The "s" needs to be removed as it is a typo.
| Affirmative with Comment | V2M2 6.3.3 | The sentences appear to read as statements that may or may not be true in fact. If these are supposed to be requirements of the AB, then start the sentence in each clause as: "The AB shall reinstate …" (no s after reinstate) | Persuasive | Requirements need the word "shall" for clairity. The "s" needs to be removed as it is a typo.
| Negative with Comment | V2M2 6.3.3 | Correction to use “shall” 6.3.3 The AB shall reinstates the accreditation of an environmental laboratory whose accreditation is revoked as specified in section 6.2, when the laboratory meets the requirements for initial accreditation, as specified in Volume 1 Module 1 of this Standard and the Primary AB’s requirements for reaccreditation. | Persuasive | Requirements need the word "shall" for clairity. The "s" needs to be removed as it is a typo.
| Negative with Comment | V2M2 General | • The new structure of V2M2 does not provide instruction for the AB to evaluate PT studies by method or technology, and therefore, how to score PT performance when the laboratory analyzes PTs based on method or technology or both. Please see my comment to V1M1: 4.3.4. • V2M2 does not include any requirements for a laboratory to report supplemental PTs within 45 days of the shipment date. In addition, it doesn’t say that scheduled PT studies can only be open for 45 days. If this requirement is added to V1, then it doesn’t need to be added to V2. | Non-Persuasive | Study open and close requirments are listed in Vol 3 for the PTP's. The laboratories are required to report prior to the close date of the PT study. PT study is defined as either a regularly scheduled PT or a supplemental PT study.
| Negative with Comment | V2M2 4.1.5 | 4.1.5 causes concern: The AB may deem a change the outcome of a PT due to lab not handling the sample as a routine sample, yet the lab must handle the sample differently if the lab may not report less than LOQ if the LOQ is higher than the PTRL. This section may further leave laboratories vulnerable to capricious interpretations by an accrediting body [for example, "etc." in c) NOTE and e)]. Suggest rewording this section to make the intended rationales and outcome clear. | Non-Persuasive | The section refers to requirements listed in the standard. This would include the V1M1 section 4.24 which allows the laboratories latitude for the analysis of PT's.
| Negative with Comment | V2M2 4.1.5c | Note:) The statement that requires labs to "handle PT samples in the same manner as a real environmental sample" is contradictory to what would be allowed by Volume 1 Module 1. The allowances in 4.2.4 of Volume 1 Module 1 would go against the standard o | Non-Persuasive | The section has been edted to refer to requirements listed in the standard. This would include the V1M1 section 4.24 which allows the laboratories latitude for the analysis of PT's.
| Negative with Comment | V2M2 4.1.5c | Note: The statement that requires labs to "handle PT samples in the same manner as a real environmental sample" is contradictory to what would be allowed by Volume 1 Module 1. The allowances in 4.2.4 of Volume 1 Module 1 would go against the standard operating procedure for a real environmental sample. | Non-Persuasive | The section has been edted to refer to requirements listed in the standard. This would include the V1M1 section 4.24 which allows the laboratories latitude for the analysis of PT's.
| Affirmative with Comment | V2M2 4.1.5c | The Note offers examples of situations where a laboratory’s “Acceptable” PT study scores might still be evaluated by an AB as unsuccessful. The example that reads “…failure to release the results to the AB before the close of the study….” is unclear to me. Who is it that failed to act in this example, the laboratory or the PT Provider? If it is the laboratory’s failure to authorize release of the study information to the AB, that is not clear. Please explain who the responsible parties are in each of the examples, as suggested here: "Note: 'Acceptable' PT study scores from a PT Provider do not automatically result in a successful evaluation of a PT study by an AB. For example, a laboratory’s failure to report an analytical method or reporting of an incorrect method, failure by the laboratory to release the results to the AB before the close of the study….” | Persuasive | “NOTE: “Acceptable” PT study scores from a PT Provider do not automatically result in a successful evaluation of a PT study by an AB. For example, failure to report an analytical method or reporting of an incorrect method, failure to provide the PTP with a release of results to the AB before the close of the study, failure to report results to the PT Provider before the closing date, failure to handle PT study samples in the same manner as real environmental samples, etc. may be cause for an unsuccessful evaluation by an AB.
| Affirmative with Comment | V2M2 4.1.5c | “NOTE: “Acceptable” PT study scores from a PT Provider do not automatically result in a successful evaluation of a PT study by an AB. For example, failure to report an analytical method or reporting of an incorrect method, failure to release the results to the AB before the close of the study, failure to report results to the PT Provider before the closing date, failure to handle PT study samples in the same manner as real environmental samples, etc. may be cause for an unsuccessful evaluation by an AB. “ Comment 2: Regarding the list of examples in the note above, I do not understand “…failure to release the results to the AB before the close of the study.” What does this mean? Is it referring to the requirement in the V1M1 VDS, Section 4.3.2, in which the laboratory is to direct the PTP to report the PT study results directly to the selected AB(s)? If so, suggest rewording the example to “…failure to direct (before the closing date) the PT Provider to release PT study performance results to the AB upon closure of the PT study…” Comment 3: The example “failure to report results to the PT Provider before the closing date” would not result in an “Acceptable” PT study score, and should not be included in the list of examples. | Persuasive | “NOTE: “Acceptable” PT study scores from a PT Provider do not automatically result in a successful evaluation of a PT study by an AB. For example, failure to report an analytical method or reporting of an incorrect method, failure to provide the PTP with a release of results to the AB before the close of the study, failure to report results to the PT Provider before the closing date, failure to handle PT study samples in the same manner as real environmental samples, etc. may be cause for an unsuccessful evaluation by an AB.
| Negative with Comment | V2M2 4.1.5e | Evaluate that the laboratory performed and implemented corrective action for failed PT Studies; Comment: What constitutes failure and if these studies are based on statistics and a 99 % level then there should be outliers for all labs. Do all outliers need to have corrective action? | Non-Persuasive | PTP's have criteria to remove outliers in the calculation of acceptance criteria generated from the study data and the PTPEC FoPT subcommittee also has outlier criteria in the calculation of the acceptance criteria.
| Negative with Comment | V2M2 4.1.5f | maintain the current accreditation status of laboratories in their program in the National Database, when the database is established Comment: The database is established but not all use it yet after 3 years. This notification should be a database function automated. | Persuasive | Changed 4.1.5 (f) to read "maintain the current accreditation status of laboratories in their program in the National Database, also known as the Laboratory Accreditation Management System (LAMS)."
| Affirmative with Comment | V2M2 4.1.5f | “maintain the current accreditation status of laboratories in their program in the National Database, when the database is established…” Comment: The “National Database,” a.k.a. the Laboratory Accreditation Management System (LAMS), was established in 2010. Suggest changing the language to “maintain the current accreditation status of laboratories in their program in the National Database, also known as the Laboratory Accreditation Management System (LAMS).” | Persuasive | Changed 4.1.5 (f) to read "maintain the current accreditation status of laboratories in their program in the National Database, also known as the Laboratory Accreditation Management System (LAMS)."
| Negative with Comment | V2M2 4.1.5g | notify all Secondary ABs of revocation of accreditation of any laboratory in their program. Comment: The database is established but not all use it yet after 3 years. This notification should be a database function automated. | Persuasive | The PTEC will forward this comment to the IT committee for consideration to add fiucntionality to the LAMS database.
| Negative with Comment | V2M2 4.1.5g | It is not always possible for a Primary AB to know where its accredited laboratories have obtained secondary accreditation. Preferably strike this clause altogether or alternately add “when the Secondary ABs are known by the Primary AB.” 4.1.5 g) notify all Secondary ABs of revocation of accreditation of any laboratory in their program. | Non-Persuasive | The requirement was reverted to the original 2009 requirement to notify secondary AB's when a laboratory accreditation has been revoked.
| Affirmative with Comment | V2M2 4.1.5g | “notify all Secondary ABs of revocation of accreditation of any laboratory in their program” Comment: Should they also not notify Secondary ABs of suspension of accreditation? Suggest changing the requirement to state “notify all Secondary ABs of suspension or revocation of accreditation of any laboratory in their program” | Non-Persuasive | The requirement was reverted to the original 2009 requirement to notify secondary AB's when a laboratory accreditation has been revoked.
| Affirmative with Comment | V2M2 4.1.5g | Is this possible? Do all Primary ABs know who all of the laboratory's Secondary ABs are? Is there a requirement for Primary ABs to collect this information from labs or for labs to submit this information to Primary ABs? | Non-Persuasive | The requirement was reverted to the original 2009 requirement to notify secondary AB's when a laboratory accreditation has been revoked.
| Affirmative with Comment | V2M2 6.1.1 | These standards specify something that the AB "may" do. If this is the case, then why is there a requirements to suspend or revoke lab accreditation for failed PTs at all? Is it the intent to require all ABs to specify penalties in their prevailing regulations or quality systems, rather than in the TNI Standards? Change the auxiliary verb to "shall" in all instances as follows: “6.1.1 The AB shall suspend a laboratory's accreditation …" | Non-Persuasive | Each AB has its own set of rules and regulations that govern the suspension or revocation of the accreditation status of a laboratroy. The word "may" was added after long debate and input from the AB's.
| Affirmative with Comment | V2M2 6.2.1 | These standards specify something that the AB "may" do. If this is the case, then why is there a requirements to suspend or revoke lab accreditation for failed PTs at all? Is it the intent to require all ABs to specify penalties in their prevailing regulations or quality systems, rather than in the TNI Standards? Change the auxiliary verb to "shall" in all instances as follows: "6.2.1 The AB shall revoke a laboratory's accreditation …" | Non-Persuasive | Each AB has its own set of rules and regulations that govern the suspension or revocation of the accreditation status of a laboratroy. The word "may" was added after long debate and input from the AB's.
| Affirmative with Comment | V1M1 4.3.7 | I have some difficulty with this section, specifically the below statement. "Note: In the case where the laboratory LOQ is greater than the PTRL: If the laboratory chooses to report a value of < LOQ and the analyte is present above the PTRL, the result will be scored as “Not Acceptable” by the PT provider." In our laboratory we analyze some metals by both ICP and GFAA, and thus have to maintain accreditation for both techniques. We have had several occasions where our calculated ICP LOQ is greater than the PTRL. I don't think results should be rated as unacceptable just because a lab's analysis range does not match with the PT supplier's analysis range. I have discussed this issue with our AB, and they are aware of the issue. One suggestion for correction was to change to an instrumental technique with greater range such as an ICP-MS -- obviously a very expensive alternative, especially for a small laboratory. | Non-Persuasive | The Committee had extensive discussions concerning this issue and decided to provide flexibility in 4.3.7 (including the note in 4.3.7) as an option to any laboratory that did not want to change their normal procedures.
| Negative with Comment | V1M1 4.3.7 | Reporting to an LOQ set by a table generated by a TNI commitee is not appropriate. Reporting values should always be based on sound scientific reasoning. Each laboratory will only be able to report to a value to which their instrument or process is capable. In many cases, it is not necessary to see below this value due to permit expectations or needs of the clients we serve. Re-establishing calibration curve ranges to incorporate a lower concentration for two to four samples per year is not only costly, in terms of time and resources, due to the method validation involved, but could lead to the need to the necessity of purchasing new equipment which would not otherwise be needed. If accepted, this draft of the Proficiency Testing module will set a precedent for an unnecessary financial burden on laboratories solely due to the TNI Standard. It will also not provide any improvement to the laboratories that must comply with its mandates. | Non-Persuasive | The Committee had extensive discussions concerning this issue and decided to provide flexibility in 4.3.7 (including the note in 4.3.7) as an option to any laboratory that did not want to change their normal procedures.
| Affirmative with Comment | V1M1 4.3.6 | Discussion on the use of "Primary" in reference to this section. | Non-Persuasive | The decision was to specify the Primary AB in section 4.4.1 and delete 4.3.6.
| Affirmative with Comment | V1M1 5.1.1b | Discussion on the use of "Primary" in reference to this section. | Non-Persuasive | The decision was to specify the Primary AB in this section.
| Affirmative with Comment | V1M1 5.1.2 | Discussion on the use of "Primary" in reference to this section. | Non-Persuasive | The decision was to specify the Primary AB in this section.
| Affirmative with Comment | V1M1 6.1.1 | Discussion on the use of "Primary" in reference to this section. | Non-Persuasive | The decision was to specify the Primary AB in this section.
| Affirmative with Comment | V2M2 5.1.1, 5.1.2, 5.1.3 and 5.1.4 | Discussion on the use of "Primary" in reference to this section. | Persuasive | The decision was to specify any AB in these sections.
| Affirmative with Comment | V2M2 | Discussion on the use of "Primary" in reference to this section. | Non-Persuasive | The decision was to specify the Primary AB in this section.