## Comments
Voting Draft Standard EL-V1M4, Clauses 1.5.2.1.1; 1.5.2.1.3; 1.5.2.2; 1.5.2.2.1; 1.5.2.2.2. Response to Comments, October 2017.
Votes cast: Affirmative 107; Affirmative with comment 8; Negative with Comment 5; Absentions 5
Disclaimer: The NELAC Institute (TNI) accepts no liability for the content of any comment on a standard. Any views or opinions on a standard are solely those of the commenter and do not
necessarily reflect those of TNI.
Vote | Section/Clause | Comment | Committee Action | Committee Comment
Affirmative with comment | 1.5.2.2.2 (b) | "Any samples analyzed in a batch associated with a failing LOQ verification shall be reanalyzed or reported with qualifiers". This statement can be accepted except for drinking water whereby EPA office of water publically stated data is not be qualified for compliance reporting. Oklahoma will abide by the EPA office of water with not accepting data with qualifications. | non-persuasive | The most stringent rules shall always apply. There are many instances in which state regulations or client requirements are more stringent than what is in the standard.
Negative with comment | 1.5.2.2.2a | I am fine with all the changes except this one. In my opinion the lab's normal acceptance ranges for recovery should not be applicable because by definition concentrations below the Reporting Limit have significantly higher variation and therefore will fall outside acceptance limits derived from higher concentrations. | editorial | The standard only states that the lab shall determine the acceptance criteria. Suggestions on establishing limits will be added to the guidance document. The committee cannot establish limits due to the wide variety of methods this standard covers. An editorial change has been made to clarify the use of LOQ specific acceptance criteria.
Affirmative with comment | 1.5.2.1.1 | Any defintion of DL should allow the use of the lowest standard on the curve as the DL. | non-persuasive | The requirements of DL in the standard currently reflect the need for any dection limit determinations to include all preparatory steps. The lowest standard of a calibration curve is not always suitable as a detection limit.
Affirmative with comment | 1.5.2.2 d and 1.5.2.2.1 c2 | If there not 20 points available or established yet for the method being performed for an LOQ, would the arbitrary 70-130% be applied as is done per SW-846 or is there additional guidance that can be provided to make these sections clear for all laboratories. | non-persuasive | If the method provides guidance limits for a low level spike then those should be used. If the method does not provide guidance limits then the lab should establish limits. The committee will ensure that guidance for this is provided in the LOD/LOQ guidance document; however guidance can not be added to the standard given the wide range of methods the standard covers.
Affirmative with comment | 1.5.2.2 | Limit of Quantitation has always been based on the lowest Standard in your calibration. Not something that has been digested or extracted. | non-persuasive | The 2009 standard states in V1M4 section 1.5.2.2 a that, "All sample-processing and analysis steps of the analytical method shall be included in the determination of the LOQ." The procedures for both the detection limit and the limit of quanitation now clarify the requirement that the determinations take into account the entire method as opposed to only the analytical method.
Affirmative with comment | 1.5.2.1.1 and 1.5.2.1.2 | Section 1.5.2.1.1: 40 CFR Part 136 Appendix B states "a minimum of two spiked samples and two method blank samples prepared and analyzed on different calendar dates is required for each instrument". The Draft Standard only requires only one spike and routine blank be analyzed on each instrument. Section 1.5.2.1.2: 40 CFR Part 136 Appendix B states "during any quarter in which samples are being analyzed, prepare and analyze a minimum of two spiked samples on each instrument, in separate batches...". The Draft Standard only requires a minimum of one verification spike and one blank analyzed on each instrument. | non-persuasive | The standard is less stringent than the current 40 CFR part 136 Appendix B. As with any requirement in which a state or federal requirement is more stringent the most stringent requirement must be followed.
Affirmative with comment | 1.5.2.2.1.a | The use of the term "spike blank" is inconsistent with "low level spike" used in 1.5.2.1.1, though they appear to refer to the same activity. Is it possible to use one or the other? | editorial | The committee agrees to the ediorial change of replacing the term "spike blank" with "low level spike" where applicable to keep the language consistent.
Affirmative with comment | 1.5.2.2.2 b | Editorial only - It appears you deleted the period at the end of the first sentence and double spaced, which is inconsistent with the rest of the section, between the second and third. It should only be single spaced. Editorial only. | editorial | Editorial change made as recommended. All puncutation, syntax and formatting are checked as the standard is finalized.
Negative with comment | 1.5.2.1.1, 1.5.2.2 and 1.5.2.2.1 | The preservative requirement that was removed from the MDL procedure (1.5.2.1.1) but was not removed from the LOQ procedure (1.5.2.2 and 1.5.2.2.1) . There needs to be consistency between both DL & LOQ; and the published methods. | editorial | This was an editorial oversight by the committee. The committee voted on October 4th to remove the preservation requirement completely from both sections of the standard. This editorial error has been corrected.
Affirmative with comment | 1.5.2.2.2 | 1.5.2.2.2 "and document a technically valid reason for the corrective action" - It does make sense that an ongoing LOQ failure suggests the need for corrective action, however documenting corrective actions (assuming this means a full blown CAPA investigation) may pose an added burden on the laboratory without much added value especially where there are poor performing analytes involved. | non-persuasive | It was not the committees intent to require a full corrective action report. The laboratory should determine the extent of correction action required. The definition of corrective action according to the TNI glossary of terms is "An action to determine and eliminate the root cause(s) of a nonconformity to prevent further recurrence of the issue." A list of acceptable corrective actions is listed in section 1.5.2.2.2 b).
Affirmative with comment | 1.5.2.2.2 | 1.5.2.2.2 "Any samples analyzed in a batch associated with a failing LOQ verification shall be reanalyzed or reported with qualifiers." - This only encourages the laboratory to do all verifications separate from samples which is quite inefficient. If all method required QC are passing, there is no added value in invalidating samples run with failing LOQ verification unless a LOQ check is required QC. | non-persuasive | It is the laboratory's responsibility to notify the client if their samples were in a batch in which a QC failure occurred.
Negative with comment | 1.5.2.1.1, 1.5.2.2, 1.5.2.2.1 | There is a contradiction between section 1.5.2.1 and 1.5.2.2. The preservative requirement was removed from the MDL procedure (1.5.2.1.1) but was not removed from the LOQ procedure (1.5.2.2 and 1.5.2.2.1). If left as is, this will likely cause issues for both the lab and accreditors. For example, it could create significant comparability issues between data for the DL and the LOQ verification, depending on the method. We need to be consistent in the requirement. | editorial | This was an editorial oversight by the committee. The committee voted on October 4th to remove the preservation requirement completely from both sections of the standard. This editorial error has been corrected.
Negative with comment | 1.5.2.1.1.b | 1.5.2.1.1.b This section references the DL determination and that the determination must incorporate the "entire analytical process". Does the "analytical process" include sample preparation? This section should be reworded similiar to 1.5.2.2.b. which includes sample preservation, processing and analysis. | non-persuasive | Your comment is noted. Thhe use of analytical process is intended to include sample preparation as presented in the current 40 CFR part 136 Appendix B. Sample preservation has been removed from the requirements listed in 1.5.2.2 b. This is also consistent with the current 40 CRR part 136 Appendix B procedure.
Negative with comment | 1.5.2.1.1.c | 1.5.2.1.1.c Removal of the word "sample" makes the sentence unclear. Recommend changing "at least one spiked" to, "at least one low level spike". Section also does not clarify what is a "low level" concentration. | editorial | Editorial change made as recommended. See the current 40 CFR part 136 Appendix B for a fuller explanation of what is a "low level" concentration.
Negative with comment | 1.5.2.1.3 | 1.5.2.1.3 If the determined DL is greater than the LOQ value, the laboratory should evaluate the DL and LOQ data and redetermine the DL or the LOQ value, not just raise the LOQ value. | non-persuasive | Although the LOQ can be raised it will be evaluated with ongoing verification analysis.
Negative with comment | 1.5.2.2.1.c.ii | 1.5.2.2.1.c.ii The laboratory established accuracy acceptance criteria should be based on in-house limits, not just an arbitrary limit, i.e.-50%. But TNI should set a minimum criteria considering the laboratory is reporting data at this concentration. | non-persuasive | The standard only states that the lab shall determine the acceptance criteria. Suggestions on establishing limits will be added to the guidance document. The committee cannot establish limits due to the wide variety of methods this standard covers.
Negative with comment | 1.5.2.2.1.c.iii | 1.5.2.2.1.c.iii. If the LOQ is less than the DL, the laboratory should evaluate the DL and LOQ data and redetermine the DL or the LOQ value, not just raise the LOQ value. | non-persuasive | Although the LOQ can be raised it will be evaluated with ongoing verification analysis.
Negative with comment | 1.5.2 | General Comments on Sections 1.5.2: Section 1.5.2.2.c. states that the LOQ must be at or above the lowest corresponding calibration standard concentration. This should state "at or below the lowest corresponding calibration standard, not above. | non-persuasive | This would be contradictory to the requirement listed in V1M4 1.7.1.1 G: "the lowest calibration standard shall be at or below the lowest concentration for which quantitative data are to be reported without qualification;"
Negative with comment | 1.5.2.1.1 and 1.5.2.2.1 | These two need to agree or whether or not the preservation is need in both the DL and the LOQ or neither. | editorial | This was an editorial oversight by the committee. The committee voted on October 4th to remove the preservation requirement completely from both sections of the standard. This editorial error has been corrected.