## Comments to DS
Quality Systems for Microbiological Testing V1M5
Disclaimer: The NELAC Institute (TNI) accepts no liability for the content of any comment on a standard.
Any views or opinions on a standard are solely those of the commenter and do not necessarily reflect those of TNI
Comment Number | Vote & Justification Persuasive/Non-Persuasive | Editorial (Y/N) | Section/Clause | Comment | Committee Action | Date Addressed | Committee Comment | Date Submitted | Date & Method of Sumbitter Notification
# | P | Y | 1.2.3.4 | Whatever they said | What was done | When | Related info/justification for decision | When | When and how
1 | P | N | 1.7.3.1.d.ii and 1.7.3.1.d.iii | Please add the word accredited either before or after the word certified lab in the Micro Module Also I have been more specific that the certified/accredited lab be certified/accredited for the specific tests being performed by the laboratory. This statement is not a requirement.(Why is this statement made in the standard if it is not a requirement?) 1.7.3.1.d.ii and iii "Analysis may be performed by another certified laboratory.” I suggest the following: Suggested change: "Analysis must be performed by a certified/accredited laboratory for the tests being measured or requested from another laboratory.” | 1.7.3.1.d.ii The laboratory shall monitor the quality of the water for disinfectant residual, conductivity, total organic carbon, and heterotrophic bacteria plate count monthly (when in use), when maintenance is performed on the water treatment system, or at startup after a period of disuse longer than one month. Analysis may be performed by another certified laboratory. If the laboratory performs these tests internally for the purpose of reagent water quality monitoring, the laboratory does not need to be accredited for these tests. When these tests are not performed internally, subcontracted work must meet the requirements of V1M2 Section 4.5.5. 1.7.3.1.d.iii The laboratory shall monitor the quality of the water for metals (Cd, Cr, Cu, Ni, Pb, and Zn) annually. and the Bacteriological Water Quality Test (to determine presence of toxic agents or growth promoting substances) annually. An exception to performing the Bacteriological Water Quality Test shall be given to laboratories that can supply documentation to show that their water source meets the criteria, as specified by the method, for High Quality (Type I) or Medium Quality (Type II) reagent water. Analysis may be performed by another certified laboratory. Subcontracted work must meet the requirements of V1M2 Section 4.5.5. | 2023-02-24 | "Certified laboratory" is not used anywhere else in the Standard, so the language was removed. In its place is a reference to V1M2 regarding subcontracted work. Updated language includes changes made in response to Comment #4 and Comment #11 | 2021-08-10 | 9/11/23, email
2 | P | N | 1.7.3.6.b.v.a | COMMENT: V1M5 1.7.3.6.b.v.a says: "The equilibrium check shall include time required after test sample addition to re-establish equilibrium conditions under full capacity load appropriate for the intended use." REQUEST: Please remove this sentence in the next adopted revision of the standard. Determining this "time" is problematic and would be error-prone unless the lab has data loggers or electronic temperature monitoring. (Opening the door to check to see if equilibrium "has been re-established yet" disturbs the equilibrium again.) Having labs determine this value is potentially detrimental to data quality since a lab may mis-understand that if they must determine this factor that they also must also "use it". It is not acceptable to "correct" incubation times by this factor. Such correction would be detrimental to sample analysis. There is no regulatory requirement regarding what this information means or what to do with it. (The "guidance document" planned on this topic will not provide regulatory authority.) This is a requirement which adds an additional burden to laboratories and does not add value to the testing or data produced. A laboratory may opt to determine this value but should not be required by the TNI Standard to do so. | 1.7.3.6.b.v.a: The laboratory shall establish the uniformity of temperature distribution conditions in incubators and water baths prior to first use after installation or service to check for any cold spots or hot spots. The equilibrium check shall include time required after test sample addition to re-establish equilibrium conditions under full capacity load appropriate for the intended use. | 2022-05-10 | Updated to clarify checking for hot and cold spots during the uniformity testing. Removes requirement that includes time to re-establish equilibrium, as there is no required use for this information. | 2021-08-12 | 9/11/23, email
3 | NP | N | 1.7.3.1.a | Section 1.7.3.1(a) The section extensively discusses sterility checks on media, water, funnels, filters, and bottles with non-selective growth media, but there is no discussion as to the proper testing protocol for these implementations. Suggestion for improvement: Add Subsection (vi) to read as follows: “Incubation temperatures and time periods shall be as specified by the manufacturer’s instructions for the non-selective media, or 24 +/- 2 hours at 35+/- 1 degrees Celsius if not otherwise specified.” | None | 2022-05-10 | This is covered in Section 1.7.3.1 preamble and 1.7.3.1.a. | 2021-09-08 | 9/11/23, email
4 | P | N | 1.7.3.1.d.ii and 1.7.3.1.d.iii | Section 1.7.3.1(d) Subsections (ii) and (iii) mention that if the specified water quality tests are performed by an outside laboratory, that laboratory has to be “certified” (accredited?) for the tests in question. What if the laboratory performs the tests in-house? Does it have to be certified for the water quality analytes tested? Suggestion for improvement: Add a NOTE or additional requirements to Subsections (ii) and (iii) to read as follows: “NOTE: If the laboratory performs these tests itself internally, the laboratory does not need to be accredited for these tests when performed for the purpose of reagent water quality monitoring.” | 1.7.3.1.d.ii The laboratory shall monitor the quality of the water for disinfectant residual, conductivity, total organic carbon, and heterotrophic bacteria plate count monthly (when in use), when maintenance is performed on the water treatment system, or at startup after a period of disuse longer than one month. Analysis may be performed by another certified laboratory. If the laboratory performs these tests internally for the purpose of reagent water quality monitoring, the laboratory does not need to be accredited for these tests. When these tests are not performed internally, subcontracted work must meet the requirements of V1M2 Section 4.5.5. 1.7.3.1.d.iii The laboratory shall monitor the quality of the water for metals (Cd, Cr, Cu, Ni, Pb, and Zn) annually. and the Bacteriological Water Quality Test (to determine presence of toxic agents or growth promoting substances) annually. An exception to performing the Bacteriological Water Quality Test shall be given to laboratories that can supply documentation to show that their water source meets the criteria, as specified by the method, for High Quality (Type I) or Medium Quality (Type II) reagent water. Analysis may be performed by another certified laboratory. Subcontracted work must meet the requirements of V1M2 Section 4.5.5. | 2023-02-24 | It is not the intent of the Standard for the laboratory to be required to be accredited for this testing when performing it for internal reagent water quality monitoring only. Language was added to make it clear that the lab does not need to be accredited for this testing if they are performing this testing for the purpose of internal reagent water quality monitoring only. Updated language includes changes made in response to Comment #1 and Comment #11. | 2021-09-08 | 9/11/23, email
5 | P | N | 1.7.3.3 | Section 1.7.3.3 The third sentence reads that all analysts must count results on the same sample, when possible. As worded, claims can be made that it may not ever be possible. Suggestion for improvement: If it really should be required that all analysts count the same sample monthly, then delete the phrase “when possible.” | 1.7.3.3: For all methods that specify a quantitative result, duplicate counts must be performed monthly on at least one (1) positive sample for each month that the test is performed. These counts may be performed on environmental samples or quality control samples. If the laboratory has multiple analysts, all each analysts must count a sample that has also been counted by another analyst. results on the same sample, when possible, with The difference between the counts shall be no more than ten percent (10%) or corrective action shall be taken difference between the counts. In a laboratory with only one (1) analyst, the same sample shall be counted twice by the analyst, with no more than a five percent (5%) difference between the counts or corrective action shall be taken. | 2022-06-14 | Updated to clarify requirement for all analysts to complete this requirement monthly. | 2021-09-08 | 9/11/23, email
6 | NP | N | 1.7.3.6.b.i and 1.7.3.6.b.ii.a.4 | Sections 1.7.3.6(b)(i) and 1.7.3.6(b)(ii)(a)(4) The specified requirements for “verification” are somewhat vague in both instances. It could be interpreted as “yes, this is indeed a thermometer” for verification purposes. Suggestion for improvement: Replace all instances of the word “verification” with “calibration verification” in both places. | None | 2022-06-14 | Section 1.7.3.6 cites V1M2 5.5.13.1, which outlines requirements for verifications | 2021-09-08 | 9/11/23, email
7 | P | N | 1.7.3.6.b.v | Under technical requirements section 1.7.3.6.b).v. Incubators, water baths, it states: The laboratory shall establish the uniformity of temperature distribution and equilibrium conditions in incubators and water baths prior to first use after installation or service. The equilibrium check shall include time required after test sample addition to re-establish equilibrium conditions under full capacity load appropriate for the intended use. I propose taking out any reference to the equilibrium. Simply shorten the sentence to include only the temperature distribution study. The laboratory shall establish the uniformity of temperature distribution and equilibrium conditions in incubators and water baths prior to first use after installation or service. The equilibrium check shall include time required after test sample addition to re-establish equilibrium conditions under full capacity load appropriate for the intended use. | 1.7.3.6.b.v.a: The laboratory shall establish the uniformity of temperature distribution conditions in incubators and water baths prior to first use after installation or service to check for any cold spots or hot spots. The equilibrium check shall include time required after test sample addition to re-establish equilibrium conditions under full capacity load appropriate for the intended use. | 2022-05-10 | Updated to clarify checking for hot and cold spots during the uniformity testing. Removes requirement that includes time to re-establish equilibrium, as there is no required use for this information. | 2021-09-16 | 9/11/23, email
8 | P | N | 1.7.3.1.a.ii and 1.7.3.1.a.iii | V1M5 1.7.3.1.a.ii It appears this section is being brought into harmony with 1.7.3.1.a.iii and although ORELAP agrees this change will allow greater operational flexibility, we are concerned this could lead to abuse. The proposed language would allow any object to be included in the sterilization batch and used for the sterility check. Examples of inappropriate objects could include brand new glassware, anti-microbial objects, and objects with significantly less surface area than the filter funnels (e.g., small glass stopper). The intent is to make sure the funnels have been sterilized, so why are we allowing the laboratory to use a proxy that may have not been exposed to sample water? This seems like it will not add value. Please consider striking this change or defining appropriate objects. | 1.7.3.1.a.ii: The laboratory shall perform a sterility check on one (1) funnel per lot of pre-sterilized single use funnels using nonselective growth media. The laboratory shall perform a sterility check on one (1) funnel / or object representative in size and use per sterilization batch sterilized in the laboratory with non-selective growth media. 1.7.3.1.a.iii: The laboratory shall perform a sterility check on at least one (1) container for each lot of purchased, pre-sterilized sample containers with non-selective growth media. The laboratory shall perform a sterility check on one (1) container / or object representative in size and use per sterilization batch sterilized in the laboratory with non-selective growth media. | 2022-05-10 | Updated to clarify that object sterilized must be representative in size and use. | 2021-10-12 | 9/11/23, email
9 | P | N | 1.7.3.2.a | V1M5 1.7.3.2.a This seems like a change rather than a clarification. This could mean an increase in media plate consumption of close to 200% for laboratories using one filtration unit per sample (2 blanks for every 1 sample). This will likely lead to a shift from one filtration unit per sample to one filtration unit for the entire batch in order to cut down on the number of blanks. If the filtration units are sterilized (and checked), and stored appropriately, then it should not be necessary to run blanks on each one. Please consider striking this change. | 1.7.3.2.a: For filtration technique, the laboratory shall conduct method blanks per the analytical method.. The filtration series may include single or multiple filtration units, which have been sterilized prior to beginning the series The analysis may utilize a filter funnel manifold with single or multiple vacuum supply ports/positions. At a minimum, the laboratory shall include a beginning and ending blank for each filtration unit manifold port/position used. In addition, the laboratory shall insert a method blank after every ten (10) samples filtered per port/position unless the laboratory uses single-use funnel sets or sanitizes filtration units by UV light (254-nm) after sample filtration. 1.7.3.2.b: The filtration series is considered ended when more than thirty (30) minutes elapses between successive filtrations. A filtration series must include filtration units that have been sterilized prior to beginning the series. During a filtration series, filter funnels shall be rinsed with three (3) 20-30 ml portions of sterile rinse water after each sample filtration. In addition, laboratories shall insert a method blank after every ten (10) samples or sanitize filtration units by UV light (254-nm) after sample filtration. The filtration series is considered ended when more than thirty (30) minutes elapses between successive filtrations. | 2022-05-10 | Update to improve flow as well as clarity of intent that each port/position used requires start/stop blanks as well a blank after 10 samples filtered on each port/position. Will require update to the implementation guidance. | 2021-10-12 | 9/11/23, email
10 | NP | N | 1.6.1.1 | How about: For all methods that specify a quantitative result, duplicate counts must be performed by qualified analysts monthly on one (1) positive sample for each month that the test is performed. Alternatively, maybe 1.6.1.1 could say something like (could be worded better, I’m sure): An individual who performs any activity involved with preparation, analysis, and/or determination of results for samples must have constant, close supervision (as defined in the laboratory's training procedure) until a satisfactory initial DOC is completed (see Section 1.6.2).Reading the section again, it does state that all analysts recount, “when possible.” So, it gives leeway to the “all analysts” bit. I had to remove markup to read correctly. | None | 2022-05-10 | Monthly analyst verification counts do not replace a DOC. We are addressing the leeway concern in Comment #5. | 2021-11-05 | 9/11/23, email
11 | P | N | 1.7.3.1.d.iii | There is not an accredited method that can meet the silica criteria for type I or type II water. | 1.7.3.1.d.iii The laboratory shall monitor the quality of the water for metals (Cd, Cr, Cu, Ni, Pb, and Zn) annually. and the Bacteriological Water Quality Test (to determine presence of toxic agents or growth promoting substances) annually. An exception to performing the Bacteriological Water Quality Test shall be given to laboratories that can supply documentation to show that their water source meets the criteria, as specified by the method, for High Quality (Type I) or Medium Quality (Type II) reagent water. Analysis may be performed by another certified laboratory. Subcontracted work must meet the requirements of V1M2 Section 4.5.5. | 2023-02-24 | It has been determined that a metals and bacterioloigcal water quality testing requirement is not needed in the Standard, as either it exists elsewhere or there are other measures in place to identify reagent water issues. Updated language includes changes made in response to Comment #1. | 2022-01-18 | 9/11/23, email