## R2C-Form
Radiochemistry - DRAFT Module 6
Disclaimer: The NELAC Institute (TNI) accepts no liability for the content of any comment on a standard.
Any views or opinions on a standard are solely those of the commenter and do not necessarily reflect those of TNI.
Comment Number | Vote & Justification (Persuasive/non-persuasive) | Editorial (Y/N) | Section/Clause | Comment | Committee Action | Date Addressed | Committee Comment | Date Submitted | Date & Method of Submitter Notification
1 | Non-Persuasive | N | 1.5.3(c) | Section 1.5.3(c) uses the phrase “entire measurement system.” Presumably, this would include all sample preparation and analytical steps. | None | Discussed 7/27/22; Voted on 8/24/22 | Yes, Section 1.5.3(c) is a subset of 1.5 (“Method Validation”) which includes all preparation and analysis steps. | 2022-01-28 | 12/28/22, email
2 | Non-Persuasive | N | 1.7.1.1(a) | Section 1.7.1.1(a) uses the phrase “radiation measurement system.” I am not sure that the “system” would pertain to one particular analytical instrument, one sample-detector combination, or all instruments of a given measurement technique or technology. | None | Discussed 7/27/22; Voted on 8/24/22 | In context within 1.7.1.1(a) itself, “system” applies to “produce consistent, comparable results across multiple detectors used for a common method.” Thus, it would apply to whatever radiation measurement system is used for a particular “common” method. This would be true whether the laboratory had only one detector or many detectors associated with the system. | 2022-01-28 | 12/28/22, email
3 | Non-Persuasive | N | 1.7.1.4 | Section 1.7.1.4 uses the phrase “detection system” in several places. | None | Discussed 7/27/22; Voted on 8/24/22 | Section 1.7.1.4 is in regard to instrument performance checks (to "measure and track the stability of key detector response-related parameters over time.") As such, it is clear in the context of use that "detection system" relates to the instrument/detector, not to other variables (e.g. method/preparation). | 2022-01-28 | 12/28/22, email
4 | Non-Persuasive | N | 1.7.2.1(b) | Section 1.7.2.1(b) uses the phrase “analytical system.” (Is this the same as a detection system? Or a radiation measurement system?) | None | Discussed 7/27/22; Voted on 8/24/22 | Section 1.7.2.1(b) is a general requirement to “process batch and sample-specific QCs to provide empirical evidence that demonstrates that the analytical system is in control”. Section 1.7.2.1(c) goes on to further detail how this relates to when “sample testing is performed that involves physical or chemical processing which affects the outcome of the test” (c.i) and when “testing is performed that does not involve physical or chemical processing…” (c.ii). | 2022-01-28 | 12/28/22, email
5 | Non-Persuasive | N |  | There is a high degree of specificity in frequency for running a “subtraction background measurement” but not how often a “short-term background check” must be run (except for the liquid scintillation detector). Again, this is an existing standard and assessing to it may not be consistent. I guess, at a minimum, the short-term background checks need to be at least as frequent as subtraction background measurements. | None | Discussed 7/27/22; Voted on 8/24/22 | Except for the case of LSC, the Standard leaves the frequency to be defined and documented the laboratory (1.7.1.6.a.i). The risk the laboratory takes by “choosing” a longer duration between short-term background checks is the potential of having to initial corrective action on a large number of samples, possibly leading to qualification or rejection of data (1.7.1.6.c). | 2022-01-28 | 12/28/22, email
6 | Non-Persuasive | N | 1.7.2.6(c) | The Draft Standard has Section 1.7.2.6(c) subdivided into (i) through (viii), but the Excel file of Expert Committee revisions to the Standard splits (iii) in the Draft Standard into (iii) and (iv). Is this correct, and Section 1.7.2.6(c) should have (i) through (ix) now? | None | Discussed 7/27/22; Voted on 8/24/22 | This is correct, and this is how it appears in the version sent. | 2022-01-28 | 12/28/22, email
7 | Non-Persuasive | N | 1.7.1(a) | The Draft Standard has a Section 1.7.1(a) but no 1.7.1(b). Is this by design, so as to put the normative requirement in 1.7.1(a) as different from the 1.7.1 general description? | None | Discussed 7/27/22; Voted on 8/24/22 | That is correct. The portion in the first 2 paragraphs of 1.7.1 are descriptive. The portion in a) is prescriptive. | 2022-01-28 | 12/28/22, email
Note on A1: Author: